## Slide 1: Welcome to GOES-17 OTRR
Peer Stakeholder-Product Validation Review (PS-PVR) for the Provisional Maturity of GOES-17 ABI L1b Products
GOES-R Calibration Working Group, 28 November 2018

## Slide 2: Acknowledgements
Current CWG Members: F. Yu, V. Kondratovich, H. Yoo, H. Qian, Z. Wang, and S. Guo.
Former CWG Members: C. Cao, F. Padula, A. Pearlman, D. Pogorzala, R. Datla, X. Shao, M. Grotenhuis, M. Chu, B. Efremova, R. Iacovazzi, L. Zhu, J. Choi, and R. Mirza-Saliq.
External Members: T. Schmit, M. Gunshor, D. Lindsey, Steve Miller, M. Coakley, K. Mitchell.
Instrument and Ground Vendors: W. Harting, J. Van Naarden, D. Gall, L. Roop, M. Wilson.
GOES-R Program for technical, administrative, and financial support: C. Rollins, A. Reth, Z. Li, M. Weinreb, J. Fulbright, E. Kline, W. Lebair, J. Kronenwetter, M. Seybold, E. Grigsby, S. Goodman, C. Keeler, A. Krimchansky.
STAR Colleagues for feedback, web, administrative, contract, financial support: J. Daniels, AWG leads, L. Brown, S. Kalluri, etc.
And many more.
Notes: All: missed anyone?

## Slide 3: Presentation Outline
Introduction
Post-Launch Product Tests (PLPT) – what we did.
ABI L1b Product Performance – what we learned.
Assessment of Maturity – looking back.
Path to Full Validation – looking ahead.
Summary and Recommendations

## Slide 4: Presentation Outline
Introduction
Post-Launch Product Tests (PLPT)
ABI L1b Product Performance
Assessment of Maturity
Path to Full Validation
Summary and Recommendations

## Slide 5: GOES-17 Post Launch Major Events
Event | Date
Launch | 2018-03-01
Orbit Insertion (became GOES-17) | 2018-03-12
PLT Started | 2018-03-26
Door Opened | 2018-04-13
First VNIR Image | 2018-04-13
LHP Anomaly Discovered | 2018-04-28
First IR Image | 2018-07-27
PLT Resumed | 2018-07-30
Beta PS-PVR | 2018-08-27
PLT Ended | 2018-10-20
Drift Started | 2018-10-24
Drift Ended | 2018-11-13
Provisional PS-PVR | 2018-11-28

## Slide 6: Context
Review of whether GOES-17 ABI L1b products have reached the Provisional Maturity as defined by the GOES-R Program.
  Review of maturity, instead of verification of requirement, although requirements are often used as reference.
  Focus on L1b products, instead of the ABI instrument, although the two are often closely related.
  Based in part on previously completed Beta review and PLTs, although different methods can and have been used to characterize the L1b products.
Significantly affected by the Loop Heat Pipe Anomaly.

## Slide 7: LHP Anomaly & FPM Temperature
GOES-16
58.3K – 58.5K
GOES-17
81K – 105K

## Slide 8: Scope
Mode 3 when ABI is stable during PLPT.
  Extend to unstable period, M4 & M6, and post-drift performance when possible.
Performance:
  Geometric (INR)
  Radiometric – Solar (VNIR)
  Radiometric – Emissive (IR)
Metrics:
  Accuracy – difference from reference or “truth”
  Precision – confidence of evaluation
  Stability – variation in time, space, mode, etc.
Technical details such as linearity and dynamic range etc. are addressed only if inquired.
  Covered by PLT.
  Will complete by Full Validation.

## Slide 9: PLPT for GOES-16 ABI
Test ID | Short Title
ABI-L1b-PLPT-001 | IR Validation IASI CrIS
ABI-L1b-PLPT-002 | IR Validation CRTM
ABI-L1b-PLPT-003 | ZRDQ Performance
ABI-L1b-PLPT-004 | VNIR Validation VIIRS
ABI-L1b-PLPT-005 | VNIR Validation Sonora
ABI-L1b-PLPT-006 | VNIR Validation RVS
ABI-L1b-PLPT-007 | Monitoring
ABI-L1b-PLPT-008 | AD Converter
ABI-L1b-PLPT-009 | Low Light SNR
ABI-L1b-PLPT-010 | Image Navigation
ABI-L1b-PLPT-011 | Image Registration
ABI-L1b-PLPT-012 | Image Stitching
ABI-L1b-PLPT-013 | VNIR Validation DCC
ABI-L1b-PLPT-014 | Detector Uniformity Stability NSS
ABI-L1b-PLPT-015 | Validation SNO NSS
ABI-L1b-PLPT-016 | L2 Validation Support NSS
ABI-L1b-PLPT-017 | Lunar NSS
ABI-L1b-PLPT-018 | Detector Uniformity
ABI-L1b-PLPT-019 | Initial Solar Calibration
ABI-L1b-PLPT-020 | IR Spatial Uniformity
Test ID | Short Title
ABI-L1b-PLPT-021 | SRF Verification
ABI-L1b-PLPT-022 | OOB Response
ABI-L1b-PLPT-023 | Polarization
ABI-L1b-PLPT-024 | Cal-to-Cal Repeatability
ABI-L1b-PLPT-025 | Linearity
ABI-L1b-PLPT-026 | Ghosting
ABI-L1b-PLPT-027 | Field Campaign – Aircraft
ABI-L1b-PLPT-028 | Field Campaign – Surface
ABI-L1b-PLPT-029 | VNIR Verification – Rayleigh
ABI-L1b-PLPT-030 | VNIR Verification – Uyuni
ABI-L1b-PLPT-031 | BDS
ABI-L1b-PLPT-032 | VNIR SNR & Dynamic Range
ABI-L1b-PLPT-033 | IR Cal Stability
ABI-L1b-PLPT-034 | MTF
ABI-L1b-PLPT-035 | Coherent Noise
ABI-L1b-PLPT-036 | IR SNR & Dynamic Range
Critical tests are in bold.
Experimental tests are in italic.

## Slide 10: Revised PLPT for GOES-17 ABI
Test ID | Short Title
ABI-L1b-PLPT-001 | IR Validation
ABI-L1b-PLPT-002 | IR Validation CRTM
ABI-L1b-PLPT-003 | ZRDQ Performance
ABI-L1b-PLPT-004 | VNIR Validation
ABI-L1b-PLPT-005 | VNIR Validation Desert Sonora
ABI-L1b-PLPT-006 | VNIR Spatial Uniformity
ABI-L1b-PLPT-007 | Monitoring
ABI-L1b-PLPT-008 | AD Converter
ABI-L1b-PLPT-009 | Low Light SNR
ABI-L1b-PLPT-010 | Image Navigation
ABI-L1b-PLPT-011 | Image Registration
ABI-L1b-PLPT-012 | Image Swath Registration
ABI-L1b-PLPT-013 | VNIR Validation DCC
ABI-L1b-PLPT-014 | Detector Uniformity Stability NSS
ABI-L1b-PLPT-015 | Validation SNO NSS
ABI-L1b-PLPT-016 | L2 Validation Support NSS
ABI-L1b-PLPT-017 | Lunar NSS
ABI-L1b-PLPT-018 | Detector Uniformity
ABI-L1b-PLPT-019 | Initial Solar Calibration
ABI-L1b-PLPT-020 | IR Spatial Uniformity
Test ID | Short Title
ABI-L1b-PLPT-021 | SRF Verification
ABI-L1b-PLPT-022 | OOB Response
ABI-L1b-PLPT-023 | Polarization
ABI-L1b-PLPT-024 | Cal-to-Cal Repeatability
ABI-L1b-PLPT-025 | Linearity
ABI-L1b-PLPT-026 | Ghosting
ABI-L1b-PLPT-027 | Field Campaign – Aircraft
ABI-L1b-PLPT-028 | Field Campaign – Surface
ABI-L1b-PLPT-029 | VNIR Verification – Rayleigh
ABI-L1b-PLPT-030 | VNIR Verification Desert
ABI-L1b-PLPT-031 | BDS
ABI-L1b-PLPT-032 | VNIR SNR & Dynamic Range
ABI-L1b-PLPT-033 | IR Cal Stability
ABI-L1b-PLPT-034 | MTF
ABI-L1b-PLPT-035 | Coherent Noise
ABI-L1b-PLPT-036 | IR SNR & Dynamic Range
ABI-L1b-PLPT-037 | Stray Light During Solar Calibration
ABI-L1b-PLPT-038 | Lunar Calibration
Critical tests are in bold.
Experimental tests are in italic.

## Slide 11: Presentation Outline
Introduction
Post-Launch Product Tests (PLPT)
ABI L1b Product Performance
Assessment of Maturity
Path to Full Validation
Summary and Recommendations

## Slide 12: ABI-L1b-PLPT-001 IR Validation
Objective
Validate thermal emissive band radiances by comparison with coincident LEO IR hyperspectral measurements
Related Requirements
MRD 737, 1493, 1503, 1513, 2130
On Schedule
Conclusion
G17 ABI IR radiance bias is small and stable during “cold” FPM period (except for B16).
No strong scene radiance Tb bias for all the IR bands
No apparent diurnal variation during the cold FPM period
Latitude/Longitude dependent Tb bias can be observed
Large Tb bias during the hot PFM period
Results
Method
Inter-calibration based on spatial, temporal, and spectral collocated homogeneous scenes, per Global Space based Inter-Calibration (GSICS) algorithm.
Time difference: <7.5 minutes for Mode3 FD data and <2.5 minutes for Mode4 FD data
Viewing geometry difference: cosine VZN diff. <1%
Geo-location difference: < ABI pixel footprint size
Spectral match: LEO spectra convolved as “pseudo-pixel”
Same target: CoVs of “pseudo-pixel” radiance < 3%
Chart: IR Bias under the “Best” Conditions
| B07 | B08 | B09 | B10 | B11 | B12 | B13 | B14 | B15 | B16
NPPCrIS-G17 | 0 | 0.475 | 0.049 | -0.035 | 0 | -0.083 | -0.124 | -0.152 | 0.015 | 0.449
N20CrIS-G17 | 0 | 0.477 | 0.086 | 0.002 | 0 | -0.064 | -0.057 | -0.099 | 0.061 | 0.469
IASI-B - G17 | -0.023 | -0.102 | 0.054 | -0.029 | -0.02 | -0.118 | -0.07 | -0.081 | 0.046 | 0.413
IASI-A - G17 | -0.078 | -0.18929 | 0.020238 | -0.064163 | -0.0015954 | -0.11085 | -0.042056 | -0.067746 | 0.052315 | 0.44892
NPPCrIS-G16 | 0 | 0.234 | -0.15735 | -0.074418 | 0 | -0.05728 | -0.08174 | -0.019712 | -0.025854 | -0.16577
IASI-B - G16 | -0.021109 | -0.13552 | -0.15538 | -0.080591 | -0.08099 | -0.1187 | -0.07383 | -0.010661 | -0.016439 | -0.19574

## Slide 13: IR Bias under “Best” Condition
LEO | B07 | B08 | B09 | B10 | B11 | B12 | B13 | B14 | B15 | B16
NPPCrIS vs. G17 | - | 0.475 (0.177) | 0.049 (0.169) | -0.035 (0.162) | - | -0.083 (0.203) | -0.124 (0.211) | -0.152 (0.218) | 0.015 (0.236) | 0.449 (0.305)
N20CrIS vs. G17 | - | 0.477 (0.188) | 0.089 (0.174) | 0.002 (0.168) | - | -0.064 (0.210) | -0.057 (0.204) | -0.099 (0.225) | 0.061 (0.252) | 0.469 (0.316)
IASI –B vs. G17 | -0.023 (0.169) | -0.102 (0.205) | 0.054 (0.192) | -0.029 (0.188) | -0.02 (0.218) | -0.118 (0.249) | -0.070 (0.280) | -0.081 (0.311) | 0.046 (0.323) | 0.413 (0.350)
IASI-A vs. G17 | -0.078 (0.168) | -0.189 (0.235) | 0.020 (0.233) | -0.064 (0.228) | -0.002 (0.327) | -0.111 (0.298) | -0.042 (0.349) | -0.068 (0.378) | 0.052 (0.392) | 0.449 (0.388)
CrIS not completely cover ABI B08 SRF
Chart: G17ABI (Gain Set=1)/G16ABI vs. CrIS/IASI, Night Time
| B07 | B08 | B09 | B10 | B11 | B12 | B13 | B14 | B15 | B16
NPPCrIS-G17 | 0 | 0.475 | 0.049 | -0.035 | 0 | -0.083 | -0.124 | -0.152 | 0.015 | 0.449
N20CrIS-G17 | 0 | 0.477 | 0.086 | 0.002 | 0 | -0.064 | -0.057 | -0.099 | 0.061 | 0.469
IASI-B - G17 | -0.023 | -0.102 | 0.054 | -0.029 | -0.02 | -0.118 | -0.07 | -0.081 | 0.046 | 0.413
IASI-A - G17 | -0.078 | -0.18929 | 0.020238 | -0.064163 | -0.0015954 | -0.11085 | -0.042056 | -0.067746 | 0.052315 | 0.44892
NPPCrIS-G16 | 0 | 0.234 | -0.15735 | -0.074418 | 0 | -0.05728 | -0.08174 | -0.019712 | -0.025854 | -0.16577
IASI-B - G16 | -0.021109 | -0.13552 | -0.15538 | -0.080591 | -0.08099 | -0.1187 | -0.07383 | -0.010661 | -0.016439 | -0.19574
Overall very accurate. Comparable with G16, except for B16

## Slide 14: Scatterplot of Tb vs. Tb Bias
G17 ABI Gain Set =1,
Collocated with night-time Metop-B/IASI
Radiance bias does not seem to depend on scene Tb.

## Slide 15: Time-Series of GEO-LEO IR Inter-Calibration (best condition: Gain Set=1, Night-time Collocations)
Long-term stable

## Slide 16: Tb Bias at Gain Set=3 (10/16/2018-10/18/2018)
Chart: Tb Bias with Gain set=3 (10/16/2018-10/18/2018)
| B07 | B08 | B09 | B10 | B11 | B12 | B13 | B14 | B15 | B16
NPP/CrIS-G17 (~7:30Z) | 0 | 0 | 0.95042 | 1.9638 | 0 | 4.3842 | -0.21631 | -2.8965 | 0.45755 | 3.2695
IASI/B-G17 (15:30Z) | 0 | -1.4282 | -0.29711 | -1.2285 | -0.15565 | -3.1201 | -0.01799 | 0.98727 | -0.3543 | -0.4213
Large Tb Bias and Variation over the Hot period

## Slide 17: Diurnal Variation
Stable over the good time. Relatively large variation or no data over the warm period.

## Slide 18: NS Direction, best condition
Latitude dependent Tb bias at some IR bans

## Slide 19: EW Direction, best condition
Slight longitude dependent Tb bias at some IR bans

## Slide 20: Special Data Collections: MESO SNO
SNPP
NOAA-20
Metop-B
Metop-A
Swath-to-Swath and MESO-to-MESO scan Tb difference can be observed

## Slide 21: ABI-L1b-PLPT-003 Zone of Reduced Data Quality
Results
Objective
Quantify measurement errors that may occur in the visible and IR channels when the sun is near and within 3o - 7.5 o for ABI VINIR and 5 o - 7.5 o from the ABI line of sight for ABI B07 3.9 um channel.
Related Requirements
MRD (2120): Radiometric accuracy for ABI IR band
<2K @300K inside restricted zone
<1K @300K outside restricted zone (> 7.5o)
Methodology
Summary
Long term monitoring shows that G17 ABI CH07 stray light inside and outside ZRDQ are both within the specification and weaker than G16 ABI CH07
Long term monitoring of ABI VNIR channel stray light has been carried out. No requirements are violated for stray light at night side for ABI VNIR channel.
Steps: Difference-Radiance-Image; Coordinate Transformation; Exponential Fitting and Estimation of Maximum Stray Light
On Schedule
Inside ZRDQ for ABI CH07
Outside ZRDQ for ABI CH07
Spec @300K
Spec @300K
Long Term Monitoring of Maximum Stray Light for G17 ABI CH07

## Slide 22: ABI-L1b-PLPT-003: ZRDQ Performance
Test ID: ABI-L1b-PLPT-003
Test Name: ZRDQ Performance
Description: Quantify measurement errors that may occur in the visible and IR channels when the sun is near and within 3 ⁰ - 7.5⁰ for VINIR and 5⁰ - 7.5⁰ for 3.9 um channel from the ABI LOS.
Due to the nature of the test, PLPT-03 is limited to within +/- 40 days of an equinox. Due to limited seasonal opportunities, this test will be continued through full validation and operations as required depending on the severity and reproducibility of stray-light effects on L1b products.
Success Criterion:. At the Provisional PS-PVR, quantify the radiance error near and within the zone of reduced data quality (ZRDQ).

## Slide 23: G17 ABI PLPT 03 Test Results-CWG
Overall stray light magnitude of G17 ABI B07 weaker than G16 ABI B07
Spike stray light (partially inside the restricted zone) becomes the dominant source of stray light for B07
G17 ABI
Other Issues: significant stray light in ABI VNIR Channel (ABI1-6) around mid- night has been observed. Since it occurs around midnight , no requirements are violated for ABI VNIR bands. But, do raise some concerns.
Stray light in other IR channels is not observable.
B07 (3.9um) | Requirement (mW/m2-sr-cm-1) | Actual Max (mW/m2-sr-cm-1)
Inside R.Z. (5o-7.5o) | < 0.077 (2K@300k) | ~0.047 ±0.0046 (1.24±0.12K@300k) (Sept. 21,2018)
Outside R.Z. (>7.5o) | < 0.038 (1K@300k) | ~0.0145 ±0.0046、(0.38±0.12K@300k) (Oct. 15, 2018)

## Slide 24: Comparison of GOES-16 and GOES-17 ABI CH07 Stray Light
GOES-16
GOES-17

## Slide 25: GOES-16 and GOES-17 ABI CH07 Stray Light Comparison
GOES-16 Daily Maximum Stray Light
GOES-17 Daily Maximum Stray Light
GOES-17 CH07 stray light is weaker than GOES-16 CH07
Both are within specs
Spec @300K inside ZRDQ
Spec @300K outside ZRDQ
Inside ZRDQ
Inside ZRDQ
Outside ZRDQ
Outside ZRDQ

## Slide 26: ABI CH01 on Aug. 12, 2018
GOES-16
GOES-17
nadir at -75.2 Deg.
nadir at -89.5 Deg.

## Slide 27: GOES-16 and GOES-17 ABI VNIR Stray Light on Aug. 12, 2018
CH01
CH02
CH03
CH04
CH05
GOES-17
GOES-17 VNIR channel stray light is weaker, but of larger spread.
GOES-16

## Slide 28: Additional Leaking Path for GOES-17 ABI?
CH01
CH02
CH03
GOES-17
CH04
CH05
GOES-16

## Slide 29: G17 ABI VNIR Channel Stray Light Monitoring (Inside and Outside ZRDQ)

## Slide 30: G17 ABI VNIR Channel Stray Light Monitoring (Inside and Outside ZRDQ)

## Slide 31: ABI-L1b-PLPT-004 VNIR Validation
Objective
Validate radiance accuracy of the VNIR bands using temporally and spatially coincident SNPP VIIRS VNIR data collected at the same viewing geometry
Related Requirements
MRD 737, 2120, and 2130
On Schedule
Conclusion
G17 VNIR radiance bias is small except for B02.
B02 is ~9% brighter than expected, similar to GOES-16.
B01/B02/B03/B05 biases show small but significant dependence on scene radiance.
  All negative slope
B05 has relatively larger uncertainty (~6%) at low albedo scenes (<30%). Uncertainty reduces as scene radiance/albedo increases
Results
Method
GSICS Ray-matching method
Time difference: <5 minutes for M3 data
Cosine of viewing zenith angle difference < 1%
Time: 17:00UTC-19:00UTC
Geolocation range: 1oS < latitude < 1oN
Azimuth angle difference: < 30°
SRF adjustment
Chart: ABI vs. SNPP/VIIRS
| B01 | B02 | B03 | B04 | B05 | B06
G17 (Raymatching) | -0.66 | 8.85 | 0.96 | -3.18 | 5.32 | -2.15
G17-DCC | 0.958 | 12.163 | 1.827 | 1.041 | 2.719 | -0.6781
G16 (Raymatching) | 4.25 | 5.6 | 1.059 | -0.925 | 4.135 | -0.424

## Slide 32: Scatterplot of Scene Dependent Refl. Ratio
F-test significant at B01/B02/B03/B05

## Slide 33: Time-series of G17 vs. NPP/VIIRS Refl. Ratio
Long-term Stable, except for B06 (B05?).
Under investigation

## Slide 34: ABI-L1b-PLPT-006 VNIR Spatial Uniformity
Results
Objective
Determine and validate scan mirror angle-dependent reflectivity corrections for the VNIR bands.
Related Requirements
MRD 2120
Methodology
Summary
Lunar chasing events were successfully conducted on 07/30/2018 and 10/20/2018.
VNIR spatial uniformity is within 1% for 5 bands.
B01 has several issues:
  Structural variation within each 15 minutes of timeline on 07/30/2018 and 10/20/2018.
  “Humps” on 10/2018
  Ongoing effort to investigate root cause to the B01 irradiance variations
On Schedule
Use of GIRO model to compensate for lunar irradiance variation for the lunar images collected as it transited across the ABI field of regard
6/1/2018
07/30/2018
10/20/2018
B02
B03
B05
B06
B01
B02
B03
B04
B05
B06
B01
B04

## Slide 35: ABI-L1b-PLPT-007 Instrument Performance Monitoring
Results
CWG GOES-17 ABI IPM results can be found at the following links:
Instrument Calibration Data https://www.star.nesdis.noaa.gov/GOESCal/G17_IPM_Ops-BothCalStats_daily.php
Engineering Data https://www.star.nesdis.noaa.gov/GOESCal/G17_IPM_EngTelemStats_daily.php
Solar Calibration Data https://www.star.nesdis.noaa.gov/GOESCal/G16_ABI_Solar_Cal_static.php
Objective
Perform automated monitoring of G17 ABI instrument performance in all 16 bands and Engineering Telemetry in near-real time at the STAR designated GOES Calibration web site.
Related Requirements
MRD737 (Long-term radiance drift); MRD2120 (Solar calibration absolute accuracy); MRD-2121 and MRD-2158 (Calibration-to-calibration radiance precision); and MRD-2130 (Product monitoring).
Methodology
Summary
CWG provides automated monitoring of G17 ABI instrument performance in all 16 bands and Engineering Telemetry in near-real time.
This IPM provides valuable and timely supports for the calibration work in CWG group, helping in resolving the issues in the algorithm/LUTs update, major anomaly events, missing dataset etc.
Monitored data represents at least 95% of INST-CAL data available for processing.
On Schedule
Flow chart of ABI IPM processing

## Slide 36: GOES-17 ABI IPM Architecture
STAR Computing Storage
STAR IPM Web Pages
All Users
Public User Facility
Read Web-ready ICVS Output Artifacts
STAR Processing Servers
Read in Imported Data and Products
STAR Web Servers
STAR Processing Servers
Read from GRIDS
STAR User
Local Interface
GOES-R Instr. Cal Data Products (NSOF ESPDS-PDA)
Perform Analyses to Update Instrument Performance Time Series
Imported Data and Products
GOES-R ICVS Dataset Storage (GRIDS)
Web-ready GOES-R ICVS Plots, Images, and Analysis
Update GRIDS
Import to SCDR or Local Storage
Public Viewport
Perform Instrument Near- and Long-Term Time-Series Analysis and Generate Output Artifacts
Write Web-ready ICVS Output Artifacts
Read
from GRIDS
GOES-17 ABI IPM Architecture

## Slide 37: GOES-17 ABI IPM Data Flow
GOES-17 ABI IPM Data Flow

## Slide 38: GOES-17 ABI IPM Components
Visible Near-Infrared (VNIR) and Infrared Channel Statistics
https://www.star.nesdis.noaa.gov/GOESCal/G17_IPM_Ops-IRCalStats_daily.php
Solar Calibration Statistics
https://www.star.nesdis.noaa.gov/GOESCal/G17_ABI_Solar_Cal_static.php
Engineering Telemetry Statistics
https://www.star.nesdis.noaa.gov/GOESCal/G17_IPM_EngTelemStats_daily.php

## Slide 39: G17 IPM Instrument Monitoring Engineering
Example: ABI LWIR FPM temperatures trending
Left: corrected FPM by CWG with the original one from INST-CAL file (right), which experienced with wrong count-to-temperature coefficients before DO.07 implemented in OE on Oct 15,2018
Smaller diurnal variation after post-drift.

## Slide 40: G17 IPM Instrument Monitoring Engineering
Example: FPM temperatures trending performs well. Lower VNIR FPM(left)
Some temperatures show lower values, but in line with life-mission range and a small diurnal variation. Other temperatures keep the same patterns as before.

## Slide 41: G17 IPM Instrument Monitoring Engineering
Internal version plotting with multiple number of ICT and Optical Bench temperatures
Generally agree well the previous stage, with few variables showing lower temperatures, but in line with the life-mission period

## Slide 42: G17 IPM VNIR bands
VNIR section performs normal after post-drift.
The trending shows the consistent pattern before G17 drift
Few smaller gaps occurred. Checked that INST-CAL file was there, will double check this later
Will discuss the period with arrows in the late slides
Gain
Space look

## Slide 43: G17 IPM VNIR bands Mean Gain
Internal version with 6 bands
On 11.14.2018, several bands experienced two gain shifting. The first one is due to solar cal at 02:54UTC, but an unknown reason for the later around 15:00-16:00UTC

## Slide 44: G17 IPM VNIR bands
Band04
Band06
On 11.14.2018, several bands experienced two gain shifting. The first one is due to solar cal at 02:54UTC, but an unknown reason for the later around 15:00-16:00UTC

## Slide 45: G17 IPM IR bands
IR section performs normal after post-drift. Example of Band08
IR sections performs normal after post-drift.
Gain
Space look

## Slide 46: G17 IPM IR bands
IR section monitoring and trending function performs normal after post-drift
Left: Internal version with 10 IR bands panel
Right: Example of Band13 gain trending in different timescale. Smaller variation of gain set2?

## Slide 47: G17 IPM IR bands
Detector level IR section monitoring and trending function performs normal after post-drift
Left: Band10 gain after post-drift (1-day and 5-day, we have a 30-day plotting)
Right: Band10 NEdT after post-drift

## Slide 48: G17 IPM Solar Calibration

## Slide 49: ABI-L1b-PLPT-009 Low Light SNR
Objective
On Schedule
Conclusion
Results
Method
Characterize G17 band 2 SNR at low light 5% albedo level, validate the low light SNR for Band 2
Related Requirements for Provisional Maturity
MRD2156 - The GOES-R Space Segment shall produce Radiance product observations in low light (5% albedo) conditions in the 0.64 micron band at a 50:1 SNR
Use ABI L1b band 2 MESO data from a Mode-3 timeline with collocated MESO-1 and MESO-2 centers.
For each of the 30 MESO scenes, compute a measure of spatial variability for each pixel
For low-light pixels with low spatial variability – i.e., mainly clear-sky homogenous ocean regions - find the temporal difference between consecutive MESO scenes
Compile results from all MESO images and determine SNR statistics for 0.2% interval over the low-light for 4%-6% albedo range.
Attempted to seek the clear-sky homogenous region with criteria of the at least 1 hour stable M3 MESO2 L1B image at the day time.
Within the region selected with mainly clear-sky homogenous condition, the low-light SNR is calculated.
CWG characterized that the mean SNR at 5% albedo for G17 in this target is 40.7.

## Slide 50: ABI-L1b-PLPT-010 Image Navigation
Objective
Quantify the navigation residuals of ABI window channels. Check navigation stability. Trend performance during PLPT period and support other PLPTs as needed.
Related Requirements
MRD522 (Navigation errors not during eclipse); MRD523 (Navigation errors during eclipse); MRD538 and MRD539 (Frame-to-frame registration (FFR)).
On Schedule
Conclusion
ABI in-frame navigation residuals meet requirements but are larger than expected performance.
ABI frame-to-frame navigation residuals meet requirements and expected performance.
Navigation residuals during eclipse meet requirements.
Navigation residuals for Modes 3, 4, and 6 have comparable values
Results
The CWG GOES-R ABI Image Navigation results can be found at the following links:
SLT Noon Navigation Residuals: https://www.star.nesdis.noaa.gov/GOESCal/G17_ABI_INR_daily.php
SLT Noon FFR:
https://www.star.nesdis.noaa.gov/GOESCal/G17_ABI_INR_daily.php
Daily Monitor Residuals:
https://www.star.nesdis.noaa.gov/GOESCal/G17_ABI_INR_daily.php
Daily Monitor FFR:
https://www.star.nesdis.noaa.gov/GOESCal/G17_ABI_INR_daily.php
Daily-Averaged Residuals for 2 Weeks:
https://www.star.nesdis.noaa.gov/GOESCal/G17_ABI_INR_daily.php
Daily-Averaged FFR for 2 Weeks:
https://www.star.nesdis.noaa.gov/GOESCal/G17_ABI_INR_daily.php
Method
Landmark use for the nav. error determination

## Slide 51: ABI-L1b-PLPT-011 Image Registration
Objective
Quantify the channel-to-channel registration (CCR) of the ABI window channels. Check CCR stability. Trend performance during PLPT period and support other PLPTs as needed.
Related Requirements
MRD529 – 533 (channel-to-channel registration). No expected performance baseline results for MRD531 and MRD533.
On Schedule
Conclusion
ABI in-FPA CCR residuals meet requirements but exceed the expected performance errors.
ABI VNIR-MWIR CCR residuals are larger than required.
ABI MWIR-LWIR CCR residuals are considerably larger than the baseline.
Recommendation: Further Kalman filter tuning, more star observations, check RDP updates, especially for the channel 7.
Results
The CWG GOES-R ABI Channel-to-Channel Registration results can be found at the following links:
SLT Noon CCR:
https://www.star.nesdis.noaa.gov/GOESCal/G17_ABI_INR_daily.php
Daily Monitor CCR:
https://www.star.nesdis.noaa.gov/GOESCal/G17_ABI_INR_daily.php
Daily-Averaged CCR for 2 Weeks:
https://www.star.nesdis.noaa.gov/GOESCal/G17_ABI_INR_daily.php
Method
Method 1: compare the navigation residuals of a channel pair.
Method 2: navigate one channel with respect to another.

## Slide 52: ABI-L1b-PLPT-012 Image Swath Registration
Objective
Assess the swath-to-swath registration (SSR) and within-frame registration (WIFR) errors of the ABI window channels. Trend performance during PLPT period and support other PLPTs as needed.
Related Requirements
MRD535 (pixel-to-pixel registration error within frame (WIFR-I) – separate two navigated radiance sample); MRD536 (pixel-to-pixel registration error within frame (WIFR-II) – register two adjacent lines/swaths)
On Schedule
Conclusion
ABI WIFR-I errors of visible channels are below the baseline and expected performance. Other channel - above the baseline but decrease with RDP updates. Constant half-pixel shift between the MESO and other frames – fix is scheduled in DO08.
ABI SSR VNIR residuals meet requirements.
ABI SSR MWIR and LWIR residuals exceed requirements, especially channel 7.
Recommendation: Further Kalman filter tuning, more star observations, check RDP updates for channel 7.
Results
The CWG GOES-R ABI Swath-to-Swath Registration and WIFR results can be found at the following links:
SLT Noon Swath Residuals:
https://www.star.nesdis.noaa.gov/GOESCal/G17_ABI_INR_daily.php
Daily Monitor STD (East-West and North-South):
https://www.star.nesdis.noaa.gov/GOESCal/G17_ABI_INR_daily.php
Daily-Averaged STD for 2 weeks (East-West and North-South):
https://www.star.nesdis.noaa.gov/GOESCal/G17_ABI_INR_daily.php
Method
WIFR I error: STD of landmark residuals.
WIFR II error: Average absolute SSR.
Swath residual: Average of landmark residuals within a swath.

## Slide 53: ABI-L1b-PLPT-013 VNIR Validation DCC
Results
Objective
To provide useful information of the G17ABI radiometric calibration stability by monitoring DCC reflectance trending.
To validate the radiance accuracy of G17ABI six visible and near infrared bands using the NPP-VIIRS data as references.
This PLPT supports MRD-737 (Long-term radiance drift), MRD-2120 (Solar calibration absolute accuracy), MRD-2130 (Product monitoring).
Methodology
The DCC pixels are described as followings:
20°S < latitude < 20°N and 120°W< longitude < 60°W
Brightness temperature at Band 14 (11µm) < 205 K
Standard deviation at Band 14 (11µm) < 1 K
Solar zenith angle < 40° and viewing zenith angle < 40 °
Reflectance standard deviation at Band 2 (0.64µm) < 0.03
Avoid sun glint area & data quality flag ≤ 1 for ABI data
Apply for Spectral Band Adjustment Factor (SBAF) derived from SCIAMACHY between ABI and VIIRS, Angular Distribution Model (ADM) adjustment
Select VIIRS DCC pixel within 1 km distance from ABI DCC pixel
Summary
DCCs reflectance from both G17ABI and NPP-VIIRS are in a good agreement in terms of pattern and shape.
G17ABI DCC reflectance of Bands 1, 2, 3, 4 and 5 is brighter than NPP-VIIRS for 0.95 %, 12.1 %, 1.8 %, 1 %, and 2.7 %, respectively but the G17ABI Band 6 is slightly darker than NPP-VIIRS about 0.7 %. The G17ABI Band 2 is much brighter than NPP-VIIRS I band.
On Schedule
11/28/2018
Left: Time series of DCC reflectance from G17 ABI and NPP-VIIRS
for six VNIR bands; right: DCC reflectance ratio for six VNIR bands.

## Slide 54: ABI-L1b-PLPT-015 NSS During SNO
Results
Objective
- Confirm the detector-to-detector variability quantified in other NSS-related tests such as PLT-018 and PLPT-018 for the VNIR bands
- Validate the calibration of each detector of the six ABI VNIR bands using VIIRS as reference
Methodology
On the date and time of SNO, perform NSS at nadir
Summary
Six SNO NSS events were requested, according to PLPT_North_South_Scan_SNO_V3.1_GOES-S.docx
The last four events were reported as “conducted” in the SOE and PLT morning briefing slides.
Three of the SNO NSS events were actually executed.
  L0 data were downloaded from LZSS and ingested with GRATDAT
  Only data conducted on 09/27 and 10/03 were ingested as NSS data. Data collocated on 10/13 were Mode 3 data
Only one SNO NSS data collocated with S-NPP will be used for further analysis
On Schedule
LEO Satellite | Data | Requested Time (UTC) 1 | Event Time (UTC) reported in SOE/PLT morning briefing slides | Actually executed Events2
Metop-B | 09/04/2018 | 15:26 | None | -
NOAA-20 | 09/17/2018 | 07:22 | None | -
S-NPP | 09/27/2018 | 19:24 | 19:22 - 19:27 | M5(NSS data)
Metop-B | 10/03/2018 | 15:26 | 07:19 - 07:24 | M5 (NSS data)
NOAA-20 | 10/03/2018 | 07:21 | 15:24 – 15:29 | M5 (NSS data)
S-NPP | 10/13/2018 | 19:24 | 19:22 – 19:27 | M3 (Not NSS data)
1: Requested by F. Padula and M. Cook
2: Successfully ingested by GRATDAT as Mode25 NSS data

## Slide 55: ABI-L1b-PLPT-017 NSS Lunar
Results
Objective
Characterize ABI instrument performance using the north-south scan (NSS) of the Moon.
Related Requirements
MRD519: detector-to-detector radiance with relative accuracy in each band within1-sigma noise
Methodology
Summary
Lunar NSS was successfully conducted on 08/30/2018 and 09/25/2018.
Each ABI detector successfully scanned across the Moon
  Although some detectors had moon-contamination space looks on 09/25/2018, valid lunar NSS for each detector was collected for each detectors from the two lunar NSS events.
Ongoing effort using the lunar NSS data for MTF, blooming, OOF, and detector-to-detector uniformity assessments.
On Schedule
ABI scans across the moon with north-south scan mode
 | 08/30/2018 | 09/25/2017
Lunar phase angle | 49o - 51o | 7o-10o
NSS over the Moon? | Yes | Yes
Issue | - | Some space calibration data were contaminated with the Moon.

## Slide 56: ABI-L1b-PLPT-018 Detector Uniformity
Results
Nearly overlapped data were partially collected.
Objective
Investigate the root cause of detector-to-detector variability for selected ABI bands.
Related Requirements
MRD519: detector-to-detector radiance with relative accuracy in each band within1-sigma noise
Methodology
Collect NSS data over Deep Convective Clouds. This establish detector SRF effects on ICT calibration.
Collect NSS data over a variety of scenes.
Collect nearly overlapped MESO data over South America that include high and low humidity, high and low altitude, high and low reflectance, ocean and forest.
Analyze the above data to evaluate the effect of surface reflectance and atmospheric absorption on detector variability.
Summary
Adequate NSS and some overlapped MESO data have been collected.
Reduced urgency and priority for this test, based on GOES16 experience and GOES17 issues related to LHP anomaly.
This test will continue later.
Delayed

## Slide 57: ABI-L1b-PLPT-019 Initial Solar Calibration
Objective
On Schedule
Conclusion
Results
Method
Establish the baseline of calibration based on Solar Calibration Target (SCT, also known as Solar Diffuser SD) for ABI visible and near infrared (VNIR) bands.
Related Requirements for Provisional Maturity
Derived calibration coefficients for ABI VNIR bands at the beginning of life (BOL).
Estimate the uncertainty of the derivation.
Compare with similar measurements and derivation pre-launch.
Reconcile discrepancies and provide the supports
Equation to calculate VNIR band detector gain
Where m is the VNIR band gain for each detector. as Solar calibration Integration factor is 9, as effective VNIR channel average spectral radiance, and Q is quadratic coefficient for each detector in bands. is the average of scene count for solar calibration and space look view, reprehensively.
G17 ABI VNIR Bands Solar Calibration trending:
https://www.star.nesdis.noaa.gov/GOESCal/G17_ABI_Solar_Cal_static.php
GOES-17 ABI VNIR bands mean gains from GS and CWG
CWG characterized G17 ABI VNIR bands gain in the overall mean and detector level in mission-life, including before and after the initial three solar cal events on Jul 31,2018
The first three solar cal gain pattern and magnitude generally agree each other with difference <0.5%.
Band02 detector gains has a pop-up-down pattern due to Q term from the column, especially in the prelaunch gain.
Band03-06 detector gain variation is large in GS, while CWG gain performs well in the variation.
GS band04-06 detector gain variation experienced major four-stage features since first three solar cal.

## Slide 58: Introduction to PLPT-019
Total 34 Solar Calibrations events (ends Nov 14, 2018) have been performed for G17 VNIR bands since the three initial solar calibration on Jul 31,2018.
Currently, G17 ABI VNIR bands solar calibration was executed normally at 02:54UTC on Nov 14, 2018 after G17 drifted to the operational west position(-137.2W).
CWG conducted a preliminary assessment of G17 ABI VNIR bands solar calibration with CWG in-house solar cal processing package and evaluated the results with comparison to the processing of GS.
We focused on the variation of band mean, detector-level gain, SNR and counts, especially explored the features at different stages of G17 (e.g. first three solar calibration, yaw flip, before and post-drift).

## Slide 59
D PLT ABI VNIR Solar Calibration
Test | Date | File Type | Channels | Start Times (on band 3) | Notes | Elev. [deg]
Yaw flip | 20180521 | INST-CAL | 1-6 | 17:56:31Z | No data | 0.37
Yaw flip | 20180530 | INST-CAL | 1-6 | 17:57:29Z | Processed | 0.38
Yaw flip | 20180611 | INST-CAL | 1-6 | 17:59:37Z | Processed | 0.37
Yaw flip | 20180613 | INST-CAL | 1-6 | 18:00:03Z | Processed | 0.37
Yaw flip | 20180616 | INST-CAL | 1-6 | 18:00:41Z | Processed | 0.38
 | 20180627 | INST-CAL | 1-6 | 06:03:06Z | Processed | -0.35
BDS col Test (7 solar cals) | 20180724* | INST-CAL | 1-6 | 06:06:33Z | Processed | -1.3
PLT008 P1-1 | 20180731 | INST-CAL | 1-6 | 06:06:25Z | Processed | -0.35
PLT008 P1-1 | 20180731 | INST-CAL | 1-6 | 06:09:25Z | Processed | *
PLT008 P1-1 | 20180731 | INST-CAL | 1-6 | 06:12:25Z | Processed | *
PLT008 P2 | 20180802 | INST-CAL | 1-6 | 06:06:17Z | Processed | -0.34
PLT008 P3 | 20180804 | INST-CAL | 1-6 | 06:06:07Z | Processed | -0.35
PLT008 P4 | 20180806 | INST-CAL | 1-6 | 06:05:55Z | Processed | -0.35
PLT008 P5 | 20180808 | INST-CAL | 1-6 | 06:05:40Z | Processed | -0.34
PLT008 P6 | 20180810 | INST-CAL | 1-6 | 06:05:23Z | Processed | -0.34
PLT008 P7 | 20180812 | INST-CAL | 1-6 | 06:05:04Z | Processed | -0.34
PLT008 P8 | 20180814 | INST-CAL | 1-6 | 06:04:43Z | Processed | -0.33
PLT008 P9 | 20180816 | INST-CAL | 1-6 | 06:04:19Z | Processed | -0.33
  PLT ABI VNIR Radiometric Calibration test began on 07/31/2018
  BDS column ABI VNIR Radiometric Calibration test began on 06/24/2018
  Yaw flip began 05.16 and ended on 06/20.
  Time and angles are based on band 3 scanning.
Notes: 59
59

## Slide 60: Mean Gain Ratio of CWG/GS to Prelaunch
CWG vs GS | band 1 | band 2 | band 3 | band 4 | band 5 | band 6
Diff before 10.18 | 1.50% | <0.5% | <0.5% | 1% | 1% | 1%
DO.07 (10.15-10.24) | 1.50% | <0.5% | 0.60% | 5% | <0.2% | 3%
DO.07 Post-drift | 1.5% | <0.5% | 0.6% | 2% | <0.4% | 1%
The difference between CWG and GS improved, compared to previous two events with DO.07 in OE since 10.15.2018 with details below

## Slide 61: Results at the early stage
12 Solar Calibrations events have been performed 07/31/2018 to 08/18/2018 (on-going) with the solar cal before the PLT beginning date.
  Confirm that G17 solar cal timing is correct, following the way of G16.
The linear coefficients trend stability is band-dependent
  band 2 detector gain shows a pop-up-down pattern. Mean band 2 gain is 7.5% higher compared to prelaunch for both CWG and GS.
  band 4 gain id lower Prelaunch for GS (17%) and CWG band 4 (11%). band 4 mean GS gain is decreasing, compared to CWG (ADR754).
  band 6 has a similar decreasing pattern as band 4.
  band 1 GS gain is 1.5% than CWG after 06.11.(ADR753)
  ADR752: GS band 3/04/05/06 detector variation is much larger than the ones of CWG and Prelaunch.
We expect some GS processing issues. The GS anomalies need to be understood and resolved.
 | Ch1 |  | Ch2 |  | Ch3 |  | Ch4 |  | Ch5 |  | Ch6 | 
Date | CWG | GS | CWG | GS | CWG | GS | CWG | GS | CWG | GS | CWG | GS
2018.08.16 | -0.4% | 1.2% | 7.2% | 7.3% | -4.6% | -5.2% | -11.2% | -17.1% | 4.0% | 4.2% | 5.3% | 1.7%
Percentage is the one compared to prelaunch for CWG and GS |  |  |  |  |  |  |  |  |  |  |  | 

## Slide 62: Solar Cal Timing
Local time
Consistent with the best time CWG predicted, similar as previous events
Local time
UTC time

## Slide 63: Solar Cal Timing and Angles
Comparison of GS actual observation (black) and CWG (red) simulated the best timing, with the difference in brown line (bottom) .
G17 solar cal timing varied ~02.54UTC in yaw flip after post-drift, following the way of G16. Timing is stable, with comparison to optimum timing CWG simulated
G17 elevation/beta angles are stable, accordingly.
Local time
Elevation Angle
Beta Angle

## Slide 64
Detector Gain of Initial Three Solar Cal
On 07.31.2018

## Slide 65
Detector Gain of Initial Three Solar Cal
On 07.31.2018
VNIR bands cal detector gain variation among first three solar cal events is within 0.5%, with a slope along the detectors.

## Slide 66
Detector Gain of CWG/GS/Prelaunch band 04-06 at different stages
10.14
10.22
DO.07
11.14
DO.07
Post-drift

## Slide 67
Similar Detector Gain Variation on 07.31.2018
  The features of detector gain in band 04-06 after post-drift on 11.14. 2018 is similar to the stage of first three solar cal on 07.31.2018
  band 04-06 detector gain experienced four-stage features since first three solar cal.

## Slide 68
band 2 Detector Gain Variation
  band 2 Detector Gain Variation did not show major pattern change since the last stages for both CWG and GS
  Similar performance between CWG and GS on 11.14.2018, compared to before-drift
  Slight difference between CWG and GS 10.14 vs 10.22
10.14.2018
10.22.2018
DO.07
11.14.2018
DO.07 Post-drift

## Slide 69: Solar Cal SCT Count

## Slide 70: Solar Cal Space look Count

## Slide 71: Solar Cal Delta Count

## Slide 72: G17 VNIR Bands Mean SNR trend
CWG
Features of mean SNR from CWG/GS:
Overall, mean GS/CWG SNR during the last solar cal events are larger than specification, slightly band-dependent on worst condition.
Solar cal mean SNR on 11.14.2018 keeps consistent as before for both CWG and GS
band 1-03 CWG/GS general agree well
GS/CWG SNR mean band 4-06 began to be in line.
GS

## Slide 73: G17 VNIR Bands Detector SNR trend
band 1
band 4
Features of detector SNR from CWG/GS:
band 4(left): do see difference since 10.18.2018, including the post-drift on 11.14.2018.
band 1(right): generally agree well

## Slide 74: Summary of PLPT-019
CWG conducted in-house solar calibration for all GOES-17 solar cal events and provide a preliminary assessment of GOES-17 ABI VNIR Bands Solar Cal. Compared to prelaunch, the mean gain of GS/CWG after initial solar cal (with 20 days) is summarized as below
VNIR bands detector gain variation among first solar cal events is within 0.5%.
After G17 drift, it shows an improvement of GS GOES-17 ABI VNIR mean gain between CWG and GS, especially in bands 4-6, compared to the before-drift with DO.07.
  Both CWG and GS band 2 mea gain increased to 8% higher compared to prelaunch
  Improved, but still large band 3-06 detector variation is large in GS (e.g. old ADR752)
  GS band 4-06 detector gain variation experienced major four-stage features since first three solar cal on 07.31.2018.
  GS/CWG band 1 diffs 1.5% after 06.11 (current ADR753)
  Agreement between CWG and GS in SNR both in visible band 1-03 and NIR band 4-06 with recent changes in GS.
 | Ch1 |  | Ch2 |  | Ch3 |  | Ch4 |  | Ch5 |  | Ch6 | 
Date | CWG | GS | CWG | GS | CWG | GS | CWG | GS | CWG | GS | CWG | GS
2018.08.16 | -0.4% | 1.2% | 7.2% | 7.3% | -4.6% | -5.2% | -11.2% | -17.1% | 4.0% | 4.2% | 5.3% | 1.7%
Percentage is the one compared to prelaunch for CWG and GS |  |  |  |  |  |  |  |  |  |  |  | 

## Slide 75: ABI-L1b-PLPT-020 IR Spatial uniformity
Results
Objective
Verify spatial uniformity for the ABI IR bands
Related Requirements
MRD 737, 1493, 1503, 1513
Methodology
Time-series of space MESO images conducted above or below the Earth Poles to investigate root cause to PICA and verify the spatial uniformity after the scan mirror emissivity LUT updates.
GEO-LEO IR inter-calibration at diurnal and geolocation variations to validate spatial uniformity of the Earth scenes
Summary
Full-disk swaths of space scans were collected via the roll maneuver right before the PLT/PLPT started.
A series of space chasing events were conducted before and after the updates of scan-mirror emissivity look-up tables.
Evidences of uniformity residual exist.
On going effort to investigate and correct the inadequacy.
On Schedule
6/1/2018
Space radiance for the space chasing swaths
Calibrated space radiance for a FD swath in the roll-maneuver position

## Slide 76: ABI-L1b-PLPT-021 Spectral Response Function (SRF)
Results
Objective
Evaluate ABI SRF for IR bands
Related Requirements
MRD506, 1493, 1503, 1513, 2120, 2130
Methodology
Summary
All the IR bands have negative bias to CrIS/IASI within less than 0.3K.
These slight cold biases may be mainly due to the erroneous blackbody PRT temperature
GOES16 ABI IR SRF are generally well characterized on-ground.
Further analysis of sounding channel SRFs may be conducted after the PRT temperature correction.
References - X. Wu and F. Yu, 2013, TGRS, 51(3), 1215-1223.
- Yu, F. and X. Wu, 2013, TGRS, 51(3). 1200-1214.
Chart: G16 ABI - CrIS/IASI Tb Bias
| B07 | B08 | B09 | B10 | B11 | B12 | B13 | B14 | B15 | B16
ABI-IASI | -0.174 | -0.225 | -0.251 | -0.18 | -0.215 | -0.235 | -0.217 | -0.148 | -0.156 | -0.305
ABI-CRIS | -0.269 | -0.202 | -0.189 | -0.222 | -0.163 | -0.18 | -0.287
Mean Tb bias to CrIS (in red) and IASI (in blue)
On Schedule

## Slide 77: ABI-L1b-PLPT-022 Out-Of-Band Response
Results
Objective
Assess ABI out-of-band (OOB) response for IR bands
Related Requirements
MRD 1493, 1503, 1513, 2120, 2130
Methodology
- Examine the differences of spatially integrated ABI measurements and spectrally integrated IASI measurements as a function of selected wavelength as obtained from concurrent IASI measurements. The two should not be correlated in the absence of OOB.
- Examine the lunar NSS data. OOB, if exists, cab be validated with extra radiance when it scans space.
Summary
For a given IR band, no significant relationship can be found between its Tb bias to IASI and any other IR bands
Several bands show relatively larger correlations with the neighbor band in the FPM.
These slight OOB energy can be observed with the lunar NSS data
Correlation between B16 and IASI Tb difference versus all the other IR bands

## Slide 78: ABI-L1b-PLPT-024 Cal-to-Cal Repeatability
Results
Objective
Assess the calibration-to-calibration repeatability for ABI.
Related Requirements
MRD519 (relative accuracy of radiance product).
Summary
The cal-to-cal repeatability is significantly less than the relative accuracy requirement of the product.
The long-term drift of the cal-to-cal repeatability is negligibly small, as well as its diurnal and seasonal variation.
The assessment only covers normal repeatability, not gain jump due to various reasons (level-shift, detector saturation, etc.)
gain variation and cal-to-cal repeatability
Methodology
Track the gain change between consecutive ICT calibration and concerts it into temperature variation
Calibration-to-Calibration Stability

## Slide 79: ABI-L1b-PLPT-025 Linearity
Objective
Validate the system linearity for selected channels of ABI.
On Schedule
Conclusion
The system requirement of linearity is met for most detectors and bands.
The results overall agree with vendor’s evaluation.
Later detector BDS update should improve the linearity of a few outlier detectors.
There are still detectors with significantly large out-of-family non-linearity.
Results
Method
Reanalyze the PLT ECAL data: nonlinearity is assessed by the maximum deviation from linear fitting between detector response of SCT/ICT and integration time.
colors: detectors
The characterized non-linearity for IR detectors

## Slide 80: ABI-L1b-PLPT-026 Ghosting
Results
Objective
Analysis of lunar NSS for the out-of-field and blooming effect.
Related Requirements
None MRD, but characterization necessary for full characterization of L1B products
Summary
Methodology
Band 1
Band 6
Band 6
Band 1
Characterized the out-of-field and blooming effect for each detector of the bands
The OOF and blooming effects are very small for all the detectors.
For the VNIR bands, B1/2/3 have less samples impacted with OOF and blooming than B4/5/6
Some B6 detectors was affected with low straylight away from the Moon. This straylight can be shown with normal lunar scan image

## Slide 81: ABI-L1b-PLPT-029 Rayleigh Scattering Calibration
Objective
Using the Rayleigh scattering calibration method, find relative and absolute calibration and trending of ABI Band 1; and relative and absolute calibration and trending of ABI Band 2.
Related Requirements
MRD-737, MRD-2120 and MRD-2130
Methodology
Success Criteria | Status
Relative and absolute calibration and trending of ABI Band 1; and relative and absolute calibration and trending of ABI Band 2. | Not Conclusive - Uncertainty of climatological mean aerosol optical depth (AOD) can lead to uncertainties and/or biases of up to 12% and 34% for Ch 1 and Ch 2 results, respectively. Method needs more accurate AOD inputs to determine relative and absolute calibration and trending.
Applied the Vermote (1992) method to ABI
Calibration targets are determined from a climatology developed by Fougnie (2002) - targets include portions of Scene-0 South East Pacific Ocean, Scene-3 North Atlantic Ocean, and Scene-8 Costa Rica Dome.
Scene solar and satellite geometry filter settings adopted from Fougnie and Henri (2009).
The NASA 6SV 1.0 Beta version is used to simulate and compare to ABI Band 1 and 2 top-of-atmosphere radiances.
Full Report:
GOES-R_ABI-L1b-PLPT-029_Report_FullValMat_24FEB2018.pptx
ABI Ch 1 Obs/Sim Albedo versus Date for Pacific SE
Results
ABI Ch 2 Obs/Sim Albedo versus Date for Pacific SE
Summary

## Slide 82: ABI-L1b-PLPT-030 VNIR Validation: Deserts
Results
Objective
To monitor any changes in the solar calibration
with the day-by-day cloud-free desert reflectance.
To validate the accuracy of ABI Channels 1, 2, 3,
5, and 6 over the Uyuni Desert using VIIRS
measurements as reference.
This PLPT supports MRD-737 (Long-term radiance
drift), MRD-2120 (Solar calibration absolute accuracy),
MRD-2130 (Product monitoring).
Methodology
Reflectance from two geophysical targets of the uniform
areas is monitored for radiometric and temporal stability.
All the G17 ABI data used for inter-comparison with NPP-
VIIRS are corrected with Spectral Band Adjustment Factor
(SBAF) coefficients as follows:
RVIIRS, ABI = (RABI – SBAF_Offset) / SBAF_Slope
where RABI is the ABI reflectance and RVIIRS,ABI is the ABI
reflectance corrected with NPP-VIIRS spectral response
function. The SBAF_Slope and SBAF_Offset are the slope and
offset of the SCIAMACHY-based linear regression coefficients.
Summary
Uyuni desert is a stable target during August –
October but reflectance is getting low during early
wintertime due to precipitation.
Desert reflectance observed from the G17 ABI five VNIR
bands is reasonably accurate compared with the NPP-
VIIRS results.
Less uncertainty in the bidirectional reflectance
distribution function effect is seen.
On Schedule
11/28/2018
Time series of reflectance variations over the Target 1 (left) and the Target 2 (right) from the G17ABI and NPP-VIIRS for five VNIR channels.

## Slide 83: ABI-L1b-PLPT-032 VNIR SNR and Dynamic Range
Objective
On Schedule
Conclusion
Results
Method
Determine the ABI visible and near infrared (VNIR) channels’ signal-to-noise ratio (SNR) and Dynamic Range at the beginning-of-life (BOL).
Related Requirements
MRD737 (Long-term radiance drift);
MRD1493/MRD1503 MRD1513(Product Measurement Accuracy);
MRD2120(solar reflective channels to within an absolute accuracy of 5%)
Equations to calculate SNR and DNR
(1)
(2)
Where SNR is the VNIR band signal-noise-ratio, dn is the dynamic range for the detector dn is digital count; subscript SCT and SPC denote Solar Calibration Target (a solar diffuser) and space, respectively; the overhead bar denotes average; σ2 is variance; DNR is the dynamic range for the detector, M the band gain.
G17 SNR monitoring and trending:
https://www.star.nesdis.noaa.gov/GOESCal/G17_ABI_Solar_Cal_static.php
GOES-17 ABI VNIR SNR mean, stdv, max, min from GS
CWG characterized the features of GOES-17 ABI VNIR bands SNR in the overall mean and detector level.
VNIR bands mean SNR exceeded the specification recently. Bands 1-3 show stable while bands 4-6 show an improvement from the early stage.
Band-dependent at worst condition, especially bands 4-6.
G17 VNIR bands dynamic ranges all meet the requirements

## Slide 84: ABI-L1b-PLPT-033 IR CALIBRATION STABILITY
Results
Objective
Evaluate the detector-to-detector, swath-to-swath, channel-to-channel radiometric calibration stability for ABI IR channels.
Related Requirements
MRD519 (relative accuracy of radiance product).
Summary
Detector-to-detector repeatability meets the specification of requirement.
Swath-to-swatch repeatability meets the specification of requirement.
Channel-to-channel repeatability meets the specification of requirement.
There is no indication of serious on-orbit degradation of the repeatability.
Methodology
Reanalyze the special scan data executed for PLT and validate the PLT results.
PLT Repeatability Reanalysis Results

## Slide 85: ABI-L1b-PLPT-034 Spatial Resolution (MTF) Evaluation of ABI
Results
Objective
Evaluate the ABI spatial resolution in terms of Modulation Transfer Function (MTF). Report post-launch E/W and N/S rising/falling MTF and comparison with MRD requirements and pre-launch MTF values
Related Requirements
MRD527: Spatial resolution (MTF)
Summary
Methodology
E/W Rising MTF of ABI CH13
Lunar NSS Data on Sept. 25, 2018
E/W Rising MTF of ABI CH01
E/W rising/falling MTF calculated with ABI MESO scan of Moon
N/S rising/falling MTF calculated with ABI NSS of Moon
Successful evaluation of ABI spatial resolution with lunar MTF
All of E/W rising MTFs and majority of N/S rising and falling MTFs meet the requirement
E/W falling MTFs of ABI need lunar observation with sharp E/W falling edge
ABI CH06 and CH11-CH13 N/S rising MTFs and ABI CH11-CH13 falling MTFs performance degrade due to transition slope in space from NSS data.
EW Rising MTF for G17 ABI
On Schedule

## Slide 86: ABI-L1b-PLPT-034 Spatial Resolution (MTF) Evaluation of ABI
N/S Rising MTF
N/S Falling MTF
E/W Falling MTF (N/A)
(Need other lunar observation to evaluate)
E/W Rising MTF

## Slide 87: ABI-L1b-PLPT-035 COHERENT NOISE
Results
Objective
Evaluate the coherent noise for ABI.
Related Requirements
MRD 506 (radiometric sensitivity and dynamic range).
Summary
The coherent noise magnitude is within the system specification of requirement.
The coherent noise level has been basically stable. There is no noticeable temporal variation.
The source of the coherent noise is better understood.
For bands 2 and 5, their coherent noise is detectable from l1b image.
Methodology
Perform Fourier transform to cropped l1b images of 2017 & 2018. Compare the Fourier series coefficients with spec. to check exceedance and on-orbit change.
Cropped Band 4 Night Time Image and Its Fourier Spectrum
FFT
Band No. | FFT Max* (2017) | FFT Max (2018) | FFT Max. in SI (2017) | FFT Max. in SI (2018)
1 | 0.0005 | 0.0045 | 0.0006 | 0.0056
2 | 0.2047 | 0.2712 | 1.2907 | 1.7098
3 | 0.0154 | 0.0143 | 0.0407 | 0.0379
4 | 0.0033 | 0.0038 | 0.0461 | 0.0534
5 | 0.0103 | 0.0094 | 0.1078 | 0.0982
6 | 0.0002 | 0.0004 | 0.0082 | 0.0139
*: units - W/m2/sr/μm
Coherent Noise Characterization Results

## Slide 88: ABI-L1b-PLPT-036 IR SNR and Dynamic Range
Objective
Determine the SNR (NEdT) and dynamic range of ABI IR channels at the beginning-of-life (BOL)
On Schedule
Conclusion
The NEdT and dynamic range of IR detectors are consistent with the vendor’s evaluation.
The NEdT of IR detectors are worse than G-16 but overall meet the system requirements except for bands 12 & 16.
There are still detectors with significantly large out-of-family NEdT.
The dynamic range of IR detectors meet the system requirements and are better than G-16.
Results
Method
Evaluate the noise of IR detectors based on the NEdT saved in INST-CAL with proper adjustment.
Calculate the saturation radiance/temperature by
Typical daily trending of NEdT and saturation temperature

## Slide 89: ABI-L1b-PLPT-037 Stray Light During Solar Calibration
Objective
Characterize possible straylight impacts during solar calibration target (SCT) looks in the operational solar calibration configuration.
Related Requirements
MRD 737, 2120, and 2130
On Schedule
Conclusion
A series of solar calibration events were conducted on 08/28/2018.
Straylight from the main optical port and the open SCC can not be detected for the events conducted on 12:05Z, 15:05Z and 18:05Z.
  ~0.16% at B06 with closed SCC at 12:05Z
Elevated SCT offset count can be observed at 09:05Z with SCC open (~4-5% of nominal solar cal data). No concern with it as SCC is normally in nearly “closed” position.
TIM meeting was held on 10/02/2018. Consensus was reached that there should be no concern for the straylight and SCC “NOTCLOSED” impacts.
Results
Method
solar calibration events beyond the nominal solar cal time.
Event (HH:MM UTC) | Solar Cal Cover (SCC)
06:05 | Open SCC
09:05 | Open SCC
12:05 | Closed SCC
12:06 | Open SCC, nominal Solar Cal.
15:05 | Open SCC
18:05 | Open SCC
Delta Count (SCT-SPL) on 09:05Z
Completed

## Slide 90: Presentation Outline
Introduction
Post-Launch Product Tests (PLPT)
ABI L1b Product Performance
Assessment of Maturity
Path to Full Validation
Summary and Recommendations

## Slide 91: VNIR Noise
Chart: Signal-to-Noise Ratio (SNR) @100% (except for 0.64 Low). Higher is better.
| 0.47 | 0.64 | 0.64 99% | 0.64 Low*10 | 0.86 | 1.38 | 1.61 | 2.25
MRD | 300 | 150 | 300 | 200 | 300 | 300 | 300 | 300
Baseline | 349 | 247 | 311 | 442 | 470 | 937 | 564 | 1008
G16 Prov Min | 650.24 | 260.8 | 350 | 300 | 563.02 | 875.94 | 447.08 | 885.79
G16 Prov Mean | 1140.62 | 427.17 | 430 | 480 | 766.75 | 1057 | 606.86 | 1050.47
G17 Prov Min | 970 | 373 | 393 | 658 | 267 | 333 | 604
G17 Prov Mean | 1164 | 473 | 488.5 | 774 | 1082 | 619 | 1088
Evaluation metrics
G16/17 @Provisional
Performance vs. requirement / baseline

## Slide 92: VNIR Noise
Chart: Signal-to-Noise Ratio (SNR) @100% (except for 0.64 Low). Higher is better.
| 0.47 | 0.64 | 0.64 99% | 0.64 Low*10 | 0.86 | 1.38 | 1.61 | 2.25
MRD | 300 | 150 | 300 | 200 | 300 | 300 | 300 | 300
Baseline | 349 | 247 | 311 | 442 | 470 | 937 | 564 | 1008
G16 Prov Min | 650.24 | 260.8 | 350 | 300 | 563.02 | 875.94 | 447.08 | 885.79
G16 Prov Mean | 1140.62 | 427.17 | 430 | 480 | 766.75 | 1057 | 606.86 | 1050.47
G17 Prov Min | 970 | 373 | 393 | 658 | 267 | 333 | 604
G17 Prov Mean | 1164 | 473 | 488.5 | 774 | 1082 | 619 | 1088
Band 1-3 are comparable to G16.
B04 failed MRD. B05 & B06 failed PB.
Mean are comparable to G16, suggesting BDS issue.
PB bars are high, even for G16.

## Slide 93: VNIR Noise Stability
Post-Drift

## Slide 94: IR Noise
Chart: Noise Equivalent delta Temperature (NEdT, mK) @300K. Lower is better
| 3.9 | 6.19 | 6.95 | 7.34 | 8.5 | 9.61 | 10.3 | 11.2 | 12.3 | 13.3
MRD | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 300
Baseline | 70 | 20 | 19 | 31 | 24 | 30 | 58 | 33 | 37 | 107
G16 Prov Mean | 80 | 13 | 15 | 25 | 18 | 20 | 28 | 25 | 22 | 29
G17 Prov Max | 106.6 | 95.4 | 84 | 151.8 | 121.2 | 368 | 92.7 | 73 | 138.1 | 1107.4
G17 Prov Mean | 88.3 | 22.5 | 22.2 | 42.6 | 22.6 | 110 | 36.5 | 30.2 | 94 | 349.6
All worse than G16.
All failed PB.
6 of the 10 bands failed MRD

## Slide 95: IR Noise
Chart: Noise Equivalent delta Temperature (NEdT, mK) @300K. Lower is better
| 3.9 | 6.19 | 6.95 | 7.34 | 8.5 | 9.61 | 10.3 | 11.2 | 12.3 | 13.3
MRD | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 300
Baseline | 70 | 20 | 19 | 31 | 24 | 30 | 58 | 33 | 37 | 107
G16 Prov Mean | 80 | 13 | 15 | 25 | 18 | 20 | 28 | 25 | 22 | 29
G17 Prov Max | 106.6 | 95.4 | 84 | 151.8 | 121.2 | 368 | 92.7 | 73 | 138.1 | 1107.4
G17 Prov Mean | 88.3 | 22.5 | 22.2 | 42.6 | 22.6 | 110 | 36.5 | 30.2 | 94 | 349.6
Except for 9.61 and 13.3 channels, mean is within MRD and comparable to G16  Negative impacts on user experience may be somehow limited.

## Slide 96: IR Noise Stability
GOES-16:
Diurnal: 27.8±0.8mK
Long Term: Stable.

## Slide 97: IR Noise Stability
GOES-17:
Diurnal: 34.5±1.5mK
Long Term: Too early but no major concerns so far.

## Slide 98: Noise with Unstable FPM Temp Warm (2018-10-23)
detector gain switch

## Slide 99: Noise with Unstable FPM Temp Cool (2018-11-14)
detector gain switch

## Slide 100: VNIR Accuracy
Chart: Difference From Reference (%)
| 0.47 | 0.64 | 0.86 | 1.38 | 1.61 | 2.25
G16 Prov | 4.2 | 7.3 | 1.7 | -5 | 4.5 | -0.2
G16 Full | 4.7 | 6.7 | 1.5 | -0.5 | 3.8 | 0.5
G17 Prov | -0.66 | 8.85 | 0.42 | -3.18 | 5.45 | -2.8

## Slide 101: VNIR Accuracy Stability
Gain Variation
Good Instrument.
Unstable Ground.
Post-Drift

## Slide 102: IR Accuracy
Chart: Difference From Reference (mK)
| 3.9 | 6.19 | 6.95 | 7.34 | 8.5 | 9.61 | 10.3 | 11.2 | 12.3 | 13.3
G16 Prov | -190 | -320 | -320 | -240 | -240 | -210 | -210 | -140 | -120 | -280
G16 Full | -180 | -210 | -250 | -190 | -220 | -230 | -220 | -150 | -160 | -290
G17 Prov | -23 | -102 | 54 | -29 | -20 | -118 | -70 | -81 | 46 | 413
Dependence of SRF on FPM temperature?

## Slide 103: IR Stability
Post-Drift

## Slide 104: Tb Bias at Gain Set=3 (10/16/2018-10/18/2018)
Chart: Tb Bias with Gain set=3 (10/16/2018-10/18/2018)
| B07 | B08 | B09 | B10 | B11 | B12 | B13 | B14 | B15 | B16
NPP/CrIS-G17 (~7:30Z) | 0 | 0 | 0.95042 | 1.9638 | 0 | 4.3842 | -0.21631 | -2.8965 | 0.45755 | 3.2695
IASI/B-G17 (15:30Z) | 0 | -1.4282 | -0.29711 | -1.2285 | -0.15565 | -3.1201 | -0.01799 | 0.98727 | -0.3543 | -0.4213
Large and unstable – to be characterized and improved.

## Slide 105: Navigation Accuracy
Chart: Navigation Errors For Selected Channels Not During Eclipse (μrad)
Baseline | 10.4 | 10.1 | 10.5 | 10.3 | 10.6 | 10.4 | 11.4 | 11.3 | 11.9 | 12.5
G16 Full | 3.92 | 0.72 | 2.69 | 1.9 | 5.5 | 3.2 | 8.2 | 4.5 | 9.4 | 5.1
G17 Prov M3 | 2.9 | 3.8 | 5.9 | 4.5 | 14.7 | 12.4 | 19 | 15.4 | 20.8 | 22.3
G17 Prov M4 | 3.5 | 4 | 4.8 | 5 | 13 | 15.4 | 18.1 | 13.2 | 15.8 | 16.7
G17 Prov M6 | 3.4 | 2.9 | 6 | 3 | 11 | 9.4 | 10.1 | 12.8 | 15.4 | 16.2
Met MRD. Worse than G16 @ Full Validation (but better than G16 @ Provisional – not shown).
0.64 & 0.86 Channels: Better than Performance Baseline. M3, M4, & M6 are similar.
Other Channels: Worse than Performance Baseline. M3 is slightly worse than M4 & M6.

## Slide 106: Navigation Accuracy
Chart: Navigation Errors For Selected Channels During Eclipse (μrad)
| 3.90 EW | 3.90 NS | 10.35 EW | 10.35 NS
Baseline | 11.4 | 11.3 | 11.9 | 12.5
G16 Full | 43 | 17.3 | 52.9 | 22.8
G17 Prov M3 | 24.8 | 10.9 | 20.5 | 16.3
G17 Prov M4 | 28.6 | 8.6 | 18.1 | 8.3
G17 Prov M6 | 22.1 | 14.7 | 25.4 | 18.9
Met MRD. Better than G16 @ Full Validation.
  Major improvement
Still worse than Performance Baseline.
M3, M4, & M6 are similar.

## Slide 107: CCR Accuracy
Chart: Channel-to-Channel Registration (CCR, μrad)
| 0.64-3.90 (EW) | 0.64-3.90 (NS) | 0.86-1.61 (EW) | 0.86-1.61 (NS) | 3.90-13.30 (EW) | 3.90-13.30 (NS) | 9.61-10.35 (EW) | 9.61-10.35 (NS)
MRD | 11.2 | 11.2 | 7 | 7 | 11.2 | 11.2 | 11.2 | 11.2
Baseline | 6.6 | 5.5 | 4.9 | 4.4 | 5.8 | 4.6 | 7 | 6.8
G16 Prov | 26.1 | 31.1 | 154.3 | 37.5 | 32.8 | 58.4 | 11.6 | 18.1
G16 Full | 5.4 | 5.9 | 1.4 | 1.2 | 6.7 | 7.2 | 2.5 | 2.9
G17 Prov M3 | 15.7 | 15.8 | 5 | 3.2 | 24.6 | 24.9 | 8.9 | 8.7
G17 Prov M4 | 13.9 | 14.3 | 6 | 4.2 | 32 | 25.2 | 10.6 | 9.9
G17 Prov M6 | 7.8 | 10.4 | 3.7 | 5.4 | 21.5 | 29 | 10.1 | 7.7
Much better than G16 @ Provisional.
Met MRD for pairs on the same FPM (0.86-1.61 and 9.61-13.3).
Not quite for pairs on VNIR & MWIR.
Long way to go for MWIR & LWIR.
No significant difference among modes.
PB remains challenging for most pairs

## Slide 108: FFR Accuracy
Chart: Frame-to-Frame Registration (FFR, μrad)
| 0.64 (EW) | 0.64 (NS) | 0.86 (EW) | 0.86 (NS) | 2.25 (EW) | 2.25 (NS) | 3.9 (EW) | 3.9 (NS) | 10.35 (EW) | 10.35 (NS)
Baseline | 13.7 | 13.5 | 13.8 | 13.5 | 13.7 | 13.6 | 13.8 | 14.8 | 13.7 | 15
G16 Prov | 3.3 | 5.7 | 0.5 | 3.1 | 2.7 | 7.8 | 3.8 | 4.1 | 12.2 | 6.1
G16 Full | 0.8 | 1.6 | 0.8 | 1.4 | 2.2 | 3.2 | 2.9 | 3.5 | 3.6 | 4.2
G17 Prov M3 | 1.1 | 1.9 | 2.6 | 3.3 | 8.1 | 7.7 | 8.7 | 8.7 | 9.7 | 11.4
G17 Prov M4 | 1.6 | 1.9 | 3.3 | 3.1 | 8.4 | 7.5 | 11.6 | 10.5 | 12.9 | 14.2
G17 Prov M6 | 1.7 | 2.25 | 3 | 3.8 | 7.4 | 9 | 11.3 | 9.3 | 13.5 | 11.4
Met the MRD and the Performance Baseline.
Mode 3 seems a bit better than M4 & M6 (Mode 3 is more stable).

## Slide 109: Average (Daily and Image) SSR Accuracy
Swath residual: the average of landmark residuals within a swath.
SSR: Absolute difference of the swath residuals averaged over all swaths of the FD image.
VNIR: Met the MRD and largely the Performance Baseline.
MWIR and LWIR: Exceed the MRD, especially for the 3.9 Channel. But better than G16 @ Provisional.
Chart: Pixel-to-Pixel Reg. Error Within Frame (WIFR) - Register Two Adjacent Lines/Swaths (μrad)
| 0.64 (EW) | 0.64 (NS) | 0.86 (EW) | 0.86 (NS) | 2.25 (EW) | 2.25 (NS) | 3.9 (EW) | 3.9 (NS) | 10.35 (EW) | 10.35 (NS)
Baseline | 4.6 | 6 | 4.7 | 6.1 | 4.7 | 6.1 | 5.4 | 6.9 | 5.5 | 7.2
G16 Prov | 3 | 1.5 | 3.6 | 10 | 50 | 17.2 | 68.9 | 36 | 11.2 | 35.7
G16 Full | 1.44 | 0.9 | 4.3 | 2.3 | 12.5 | 10.1 | 12.9 | 9.7 | 18.1 | 9.1
G17 Prov M3 | 3.1 | 4.5 | 6.6 | 6.5 | 9.2 | 7.3 | 41.4 | 12.7 | 18.8 | 10.5
G17 Prov M4 | 4.3 | 4.7 | 7.4 | 6.6 | 10.4 | 7.9 | 41 | 16.3 | 21.4 | 13.6
G17 Prov M6 | 4 | 4.8 | 7.4 | 5 | 9.7 | 5.7 | 38.9 | 13 | 20.4 | 11.7

## Slide 110: INR Stability
Performance history, Sept 14 to Oct. 23, 2018.
  Residual, r-CCR, FFR, WIFR, and SSR.
  Mean and standard deviation for each.
  VNIR star comparison, PM vs. CENRAIS.
Major navigation event
  X-RDP update 9/20.
  Y-RDP update 9/21.
  Yaw flip 9/24, followed by GNC-only navigation period and Kalman filter reset.
  Mode 6 started, 9/25.
  Mode 4 started, 9/30.
  Mode 6 started, 10/1.
  RDP update, 10/3.
  Mode 3 started, 10/5.
  Diagonal Scan for VIS stars implemented, 10/15.

## Slide 111: Daily Residuals, 9/14 – 10/23
Legend: ABI daily-average residuals are plotted. Time is UTC. No VNIR residuals during nighttime. Uninterrupted data stream is joined by lines. Empty spaces are data gaps.
Discussion: Large LOS and FPM offsets (about 600 urad) were observed in the first light image. EW (x) and NS (y) navigation errors are reduced by the order of magnitude by the static LOS correction, allowing the star observations and routine navigation. Residual NS offset by about 150 urad was eliminated on September 13. The plots show the navigation error evolution since September 14. Dense timeline of testing and RDP updates has contributed to an uneven pattern. However, even at the beginning of the period the navigation residuals are comfortably low, a marked improvement with respect to ABI on GOES-16.
 -160 urad before 09/13
1
2
3
4
5
6
7
8
9
1
2
3
4
5
6
7
8
9

## Slide 112: Daily Residuals STD, 9/14 – 10/23
Legend: ABI daily-average Standard Deviations (STDs) are plotted. Time is UTC. No VNIR residuals during nighttime. Uninterrupted data stream is joined by lines. Empty spaces are data gaps.
Discussion: Large LOS and FPM offsets (about 600 urad) were observed in the first light image. EW (x) and NS (y) navigation errors are reduced by the order of magnitude by the static LOS correction, allowing the star observations and routine navigation. Residual NS offset by about 150 urad was eliminated on September 13. The plots show the navigation STD evolution since September 14. Dense timeline of testing and RDP updates has contributed to an uneven pattern. However, even at the beginning of the period the navigation STDs are comfortably low and show a tendency to decreasing, a marked improvement with respect to ABI on GOES-16.

## Slide 113: Daily r-CCR, 9/14 – 10/23
Legend: ABI daily-averaged vector Channel-to-Channel Registration (CCR) errors are plotted. Time is UTC. No VNIR residuals during nighttime. Uninterrupted data stream is joined by lines. Empty spaces are data gaps.
LEFT plot shows CCR for within-FPA channel pairs, Red color for the channel pairs from LWIR, green for MWIR, and blue for VNIR. They alternate with black in order to distinguish neighboring plots.
MIDDLE plot shows the same for the channel pairs including MWIR window channels 7 and 11. Red color for the channel pairs from LWIR, and blue for VNIR. They alternate with black in order to distinguish neighboring plots.
RIGHT plot shows the same for VNIR-LWIR pairs of channels. Red color alternates with black in order to distinguish neighboring plots.
Discussion: CCR vector error shows less navigation problems for within-FPA pairs of channels (left plot) in comparison with the channel pairs from different FPAs (middle and right plots). That shows the Kalman filter being one of the focus points of the navigation. With respect to that, co-registration of channels 2 and 7 is the key parameter.
In-FPA CCR
MWIR - VNIR & LWIR CCR
VNIR - LWIR CCR

## Slide 114: Daily r-CCR STD, 9/14 – 10/23
Legend: ABI daily-averaged vector Channel-to-Channel Registration (CCR) STDs are plotted. Time is UTC. No VNIR residuals during nighttime. Uninterrupted data stream is joined by lines. Empty spaces are data gaps.
LEFT plot shows CCR STD for within-FPA channel pairs, Red color for the channel pairs from LWIR, green for MWIR, and blue for VNIR. They alternate with black in order to distinguish neighboring plots.
MIDDLE plot shows the same for the channel pairs including MWIR window channels 7 and 11. Red color for the channel pairs from LWIR, and blue for VNIR. They alternate with black in order to distinguish neighboring plots.
RIGHT plot shows the same for VNIR-LWIR pairs of channels. Red color alternates with black in order to distinguish neighboring plots.
Discussion: CCR vector STD shows less navigation problems for within-FPA pairs of channels (left plot) in comparison with the channel pairs from different FPAs (middle and right plots). That shows the Kalman filter being one of the focus points of the navigation. With respect to that, co-registration of channels 2 and 7 is the key parameter.
VNIR - LWIR CCR STD
In-FPA CCR STD
MWIR - VNIR & LWIR CCR STD

## Slide 115: Daily FFR, 9/14 – 10/23
Legend: ABI daily-averaged Frame-to-Frame Registration (FFR) errors are plotted (squares). Daily-averaged absolute FFR values are plotted as diamonds. Time is UTC. No VNIR residuals during nighttime. Uninterrupted data stream is joined by lines. Empty spaces are data gaps.
Discussion: Daily FFRs are very close to zero, as expected. Daily absolute FFRs show increasingly-stable navigation during the period. CENRAIS was unable to process Mode 6 data for FFR except for one day, that is why plots show total zero for these days.

## Slide 116: Daily FFR STD, 9/14 – 10/23
Legend: ABI daily-averaged Frame-to-Frame Registration (FFR) STDs are plotted (squares). Daily-averaged absolute FFR values are plotted as diamonds. Time is UTC. No VNIR residuals during nighttime. Uninterrupted data stream is joined by lines. Empty spaces are data gaps.
Discussion: Daily FFR STDs react more dynamically on the changing modes and RDP LUTs, showing the same tendency to decrease by the end of the period. CENRAIS was unable to process Mode 6 data for FFR except for one day, that is why plots show total zero for these days.

## Slide 117: Daily WIFR, 9/14 – 10/23
Legend: ABI daily-averaged Within Frame Registration (WIFR) errors are plotted. Time is UTC. No VNIR residuals during nighttime. Uninterrupted data stream is joined by lines. Empty spaces are data gaps.
Discussion: WIFR is very sensitive to the image quality and shows an overall tendency to better image quality (lower WIFR), except for the last day, when a RDP LUT deployment error was made. Some small spikes correlate with the days when other deployment mistakes were made.

## Slide 118: Daily WIFR STD, 9/14 – 10/23
Legend: ABI daily-averaged Within Frame Registration (WIFR) STDs are plotted. Time is UTC. No VNIR residuals during nighttime. Uninterrupted data stream is joined by lines. Empty spaces are data gaps.
Discussion: WIFR STD is very large when the image quality is low (lower-quality images have enhanced diurnal dependence). STD also shows an overall tendency to better image quality (lower WIFR STD), except for the last day, when a RDP LUT deployment error was made.

## Slide 119: Daily SSR, 9/14 – 10/23
Legend: ABI daily-averaged Swath-to-Swath Registration (SSR) errors are plotted (squares). Daily-averaged absolute SSR values are plotted as diamonds. Time is UTC. No VNIR residuals during nighttime. Uninterrupted data stream is joined by lines. Empty spaces are data gaps.
Discussion: Daily SSRs are very close to zero, as expected (image shape and size are stable). Decrease in the absolute SSR values is desirable and remains an issue.

## Slide 120: Daily SSR STD, 9/14 – 10/23
Legend: ABI daily-averaged Swath-to-Swath Registration (SSR) STDs are plotted (squares). Daily-averaged absolute SSR STD values are plotted as diamonds. Time is UTC. No VNIR residuals during nighttime. Uninterrupted data stream is joined by lines. Empty spaces are data gaps.
Discussion: Daily SSR STDs are very close to zero, as expected (image shape and size are stable). Decrease in the absolute SSR STD values is desirable and remains an issue.

## Slide 121: VNIR Stars Comparison (PM – CENRAIS, 7/10)
Legend: CENRAIS-measured VNIR daily-averaged navigation residuals (EW and NS) are plotted in red over the GRITT plot of the VNIR star residuals. Also plotted are mean (solid) and 3-σ (dashed) of the differences.
Discussion: Good agreement is observed.
-0.07
-0.55

## Slide 122: Presentation Outline
Introduction
Post-Launch Product Tests (PLPT)
ABI L1b Product Performance
Assessment of Maturity
Path to Full Validation
Summary and Recommendations

## Slide 123: Provisional Validation
Preparation Activities | Assessment
Validation activities are ongoing and the general research community is now encouraged to participate. | Product quality has been bench marked. Comprehensive validation activities are ongoing. The general research community has been actively participating and providing feedback.
Severe algorithm anomalies are identified and under analysis. Solutions to anomalies are in development and testing. | Operational and algorithm anomalies have been identified and prioritized based on their criticality, root cause investigation and maturity of solution, implementation schedule, etc.
Incremental product improvements may still be occurring. | Incremental product improvements are expected.

## Slide 124: Provisional Validation
End State | Assessment
Product performance has been demonstrated through analysis of a small number of independent measurements obtained from select locations, periods, and associated ground truth or field campaign efforts. | This has been demonstrated.
Product analysis is sufficient to communicate product performance to users relative to expectations (Performance Baseline). | This has been communicated during PS-PVR.
Documentation of product performance exists that includes recommended remediation strategies for all anomalies and weaknesses. Any algorithm changes associated with severe anomalies have been documented, implemented, tested, and shared with the user community. | These have been documented, for example in test report and ReadMe file.
Testing has been fully documented. | Testing has been fully documented in PLPT Report.
Product is ready for operational use and for use in comprehensive cal/val activities and product optimization. | We concur.

## Slide 125: Presentation Outline
Introduction
Post-Launch Product Tests (PLPT)
ABI L1b Product Performance
Assessment of Maturity
Path to Full Validation
  Progress since Beta
  Overall assessment of current status
  Steps to reach the goal
Summary and Recommendations

## Slide 126: Progress Since Beta
Reviewed PLT final reports:
  Completed.
Performed PLPTs:
  Started and on-going.
Assisted anomaly resolution:
  Discover/Confirm; Characterize; Investigate; Correct; Implement; Verify; and Document.
Prepared preliminary assessment of findings for maturity promotion:
  Completed.
Managed risks:

## Slide 127: Risks to Reaching Provisional
Risk Name | Description | Impact | Status
Mission Data | IF the lack of accurate instrument engineering and calibration data continues; | THEN the efficiency or even the effectiveness of CWG work will be compromised. | Mostly resolved with DO.07. Very important and appreciated.
Contract | IF the problematic transition to ProTech contract continues; | THEN CWG may not be able to complete all tasks on time. | Many setbacks and complications indeed. Resolved now.
Funding | IF the FY19 fund cannot be …; | THEN CWG may not be able to complete all tasks on time. | Did not recover/resolve. Remains a risk.
Notes: Following are things important to CWG but I’m not sure I should bring it to the review:
Communication with Ground. We often do not know what happened there.
Documentation about Ground algorithm.
FY19 funds be available by June 1 (July to August in the past) to smooth out several unexpected changes associated with contract changes.

## Slide 128: Status at Provisional – GOES-17
 | Past Performance | Current Status | Future Outlook
IR Calibration | Accurate and stable when FPM temp is stable. Large but tolerable noise for most channels. Less known and may be unacceptable performance when FPM unstable. | Resolve and monitor anomalies. Extend to performance when FPM temperature is unstable. Characterize. Seek opportunity to improve. | Incremental to moderate improvement for performance when FPM temperature is stable. Need to define the performance when FPM temperature is unstable, and chart a path forward.
VNIR Calibration | High accuracy. Except for B02 bias. Low noise. Instrument is stable. Except for B03 BRDF Ground seems improving. | Resolve and monitor anomalies. | Incremental improvements. Will reach Performance Baseline.
INR | Decent performance After the correction of 160 µrad error. Frequent RDP update without adverse effects. The INR team (Flight, Ground, IPATS, CWG) are more skillful and confident in operating Kalman filter for ABI. | RDP LUT verification; Additional tuning of the Kalman filter for better temporal stability of navigation; Include more star observation in the Mode 3 and 6 timelines to increase the image quality. | Incremental improvements. Will reach Performance Baseline.

## Slide 129: Status at Provisional – GOES-16
 | Past Performance | Current Status | Future Outlook
IR Calibration | Accurate and stable. Acceptable noise, dynamic range, bias, and diurnal stability from beginning. Critical issues: None. | Performance has been good and stable without critical issues. Major remaining issues: PICA Midnight (space look shift?) | Good start Lack of hands-on experience, due to lack of INST-ENG data.
VNIR Calibration | Accurate, but unstable Critical issues: Initial bias Banding and striping 10% bias | Critical issues have been resolved. Overall performance acceptable. B02 bias well known. Major remaining issues: B02 differs from prelaunch ~7% B02 jumped ~4% Striping (nonlinearity?) Decouple solar cal and update | Team expertise. Good understanding of major issues. Good instrument. Need to work on some hard core problems that may require algorithm change
INR | Often inaccurate and unstable. Good first light navigation – good foundation without Kalman filter. Kalman filter works under favorable conditions. Critical issues: Ingest of star data, esp. IR Continuity of coregistration files Gyro spike RDP | Critical issues have been resolved. Recent performance becomes acceptable. Major remaining issues: Stability Large NS navigation accuracy (update RDP). CCR (Channel-to-Channel Registration) | Team expertise. Concept of KF-based INR strategy has been demonstrated. Good understanding of major issues. Corrections are in various stages of implementation in GS. This has been expected as the difficult problem.

## Slide 130: Path to Full Validation
Complete PLPTs.
Resolve current issues.
Monitor for new anomalies.
IR calibration when FPM temperature is not steady.
  Characterize performance and impact on L1b products.
  Seek for improvements.
Prepare comprehensive assessment of instrument performance and L1b product quality for maturity promotion.
Identify and manage risks.

## Slide 131: Risks to Full Validation
Risk Name | Description | Impact | Mitigation and Schedule
Communication with GS | IF CWG does not know what happens at GS; | THEN it becomes difficult for CWG to plan its actions. | Raise awareness. Stress the importance and urgency.
Documentation | IF CWG does not have the documents it needs; | THEN it becomes difficult for CWG to work efficiently and effectively. | Raise awareness. Stress the importance and urgency.
Funding | IF the FY19 fund cannot be dispensed by June 1, 2019 (1-2 months earlier than before); | THEN CWG work may be interrupted. | Raise awareness.
Contract / Grant | IF UMD grant renewal encounters unexpected difficulties; | THEN CWG work may be interrupted. | Watch closely and plan accordingly.

## Slide 132: Major Issues
ADR 792: B02 bias (~10% brighter). #100.
ADR 814: B03 gain oscillation, which may or may not be due to K-coeff. #101, #82, #33.
ADR ###: B16 bias, which may or may not be due to SRF dependence on FPM temperature. #102, #14.
ADR 825: B05 striping (not shown).
ADR ###: B01 variation (#34).
ADR ###: Temporal variation of IR space count (not shown).
ADR ###: Scene dependence of VNIR bias (#32).
ADR ###: Residual PICA (not shown).
ADR ###: IR space count difference in two MESO swaths (not shown).

## Slide 133
ADR # | Description | PoC | Module* | Status
254 | Positive Radiance Bias (9%) in ABI B02 | FY | G-16 | Active
292 | ABI INST-CAL Offset Field Irregularities | ZW | - | Active
300 | Large Navigation Error in ABI Meso Scan | VK | - | Active
304 | DO.04.04 ABI INST-CAL-ENG Data Show Irregularities | ZW | - | Active
352 | Bad ABI M4 Space Look Means | HQ | G-16 | Active
486 | ABI CH2 INST-CAL Data Record Number Inconsistency | HQ | G-16 | Active
500 | ABI update to CDRL 80 Volume 3 Rev F | - | G-16 | Completed
503 | ABI cold pixels near bright scenes | ZW | - | Active
583 | Missing ABI INST-CAL .XML files in LZSS | FY | - | Active
594 | ABI Lunar Intrusion Level Latch up: Longer Term Fix | ZW | G-16 | Active
616 | ABI Meso Navigation Bias relative to Full Disk/CONUS | VK | G-16 | Active
642 | Two ABI M2 files per channel for every SCT event | HQ | G-16 | Hold
704 | ABI CDRL42 Updates for GS PG L1b Processing | ZW | G-16 | Active
729 | ABI Saturated Internal Cal Target Counts | FY | G-16 | Active
732 | G16 ABI Band3 Striping | HQ | G-16 | Hold
*: If not indicated, it means the ADR is applied to all on-orbit modules
ADR Monitored by CWG

## Slide 134
ADR # | Description | PoC | Module | Status
741 | ABI update to CDRL 80 Volume 3 Rev G | HQ | G-16 | Active
753 | G17 Band01 gain raised 1.5% since 06.11.2018 | HQ | G-17 | Active
769 | same time coverage timing at G16/G17 lunar L1alpha | FY | G-17 | Active
770 | Striping in G-17 ABI IR images | ZW | G-17 | Active
771 | G-17 ABI IR detectors with large NEdT | ZW | G-17 | Active
775 | ABI/G16 Anomalous Kalman Filter Response | VK | G-18 | Hold
790 | G17 ABI B01 RVS issue | FY | G-17 | Active
791 | G16 ABI B04/B06 lunar data mis-processed | FY | G-16 | Active
792 | G17 Band02 Post-launch Gain 6.2% over Pre-launch | HQ | G-17 | Active
800 | empty data field in INST-CAL | ZW | - | Active
803 | G-17 lunar intrusion rejection algorithm revision | ZW | G-17 | Active
808 | temporal change of GOES-17 space look count | ZW | G-17 | Active
809 | G17 Periodic Infrared Calibration Anomaly Residue | FY | G-17 | Active
811 | G-17 space look spatial uniformity | ZW | G-17 | Active
814 | G17 Band03 gain variation | HQ | G-17 | Active
ADR Monitored by CWG

## Slide 135
ADR # | Description | PoC | Module | Status
815 | G17 band02 prelaunch gain variation with BDS col=2 | HQ | G-17 | Active
817 | fill-in space look statistics in INST-CAL | ZW | - | Active
825 | G-17 Band 5 Striping | HQ | G-17 | Active
826 | G-17 ICT PRT Coeff. LUT Update | ZW | G-17 | Active
831 | output gain setting info | ZW | G-17 | Active
838 | G17 VNIR bands striping since 10.15.2018 | HQ | G-17 | Active
839 | G17 Solar cal gain change since 10.18.2018 | HQ | G-17 | Active
853 | Missing GOES-17 ABI Lunar L1alpha data from LZSS | FY | G-17 | Active
856 | G16/17 Band 2 INST-CAL subset file missing | HQ | - | Submitted
TBD | GPA’s handling of saturated detectors |  | G-17 | Proposed
ADR Monitored by CWG

## Slide 136
ADR # | Description | PoC | Module | Status
574 | Corrections to ABI Image Navigation Ground Alg | VK | G-17 | Active
759 | ABI Temperature difference compared to CASSIE | ZW | - | Active
777 | Controlling deployment of ABI SCT-derived gains | HQ | - | Active
807 | ABI FD scan does not update ICT gains mid-scene | ZW | G-17 | Active
823 | G17 ABI ch11 detector 203 is bad | ZW | G-17 | Active
827 | ABI L1b temperature data quality flag (TDQF) | ZW | G-17 | Active
832 | GOES-16 X-axis Bias Anomaly | SG | G-16 | Active
858 | ABI FM1 Kalman LUT QBoost Update | ? | G-16 | Submitted
859 | ABI CDRL 79 FM2 Rev M Q tables | ? | G-17 | Submitted
ADR Monitored by CWG

## Slide 137: Presentation Outline
Introduction
Post-Launch Product Tests (PLPT)
ABI L1b Product Performance
Assessment of Maturity
Path to Full Validation
Summary and Recommendations

## Slide 138: Summary and Recommendations
Introduced the context in which this review is conducted.
  PLTs; LHP Anomaly; Revised PLPTs; Scope.
Reviewed the Post-Launch Product Tests (PLPT) activities.
Highlighted key performances of ABI L1b product.
  INR; VNIR; IR.
  Accuracy is on par with GOES-16 in general.
  Elevated radiometric noise for IR, as expected because of LHP anomaly, but the noise with stable FPM temp is largely within MRD.
  Too soon for credible assessment of stability, but no major concerns yet.
Assessed the Provisional Maturity per GOES-R Program
Outlined the path to Full Validation maturity.
Recommendation: the GOES-17 ABI L1b products be designated the Provisional Maturity per GOES-R Program.